## Slide 1: NeuroDOT
Tutorial:
Generating a Light Model

## Slide 2
NeuroDOT Set-Up - Checklist
Before starting any tutorial, please make sure you’ve done the following
Set your path so that all required folders and subfolders are included
  Folders w/subfolders to include
    Directory containing the NeuroDOT toolbox
    Any other toolbox used for whatever task
    Your personal directory
  Folders to include
    Output directory
Set your output directory so that you can save files to the desired location
  Change your current directory to your output directory
    cd(outputdir)

## Slide 3
NeuroDOT Set-Up – Output Directory
Once you’ve set your path, the next thing you should do is set your output directory.
  The output directory must be located within your path
  If it isn’t already in your path, you must add it to your path
You can set your own path to an output directory in the line shown above
An output directory specifies the location where files will be saved
  MATLAB, by default, will save files to the current directory
The following line of code will change your current directory to the output directory that you just specified 
Now you can save outputs directly to your output directory
Example output directory

## Slide 4: Diffuse Optical Tomography
In diffuse optical tomography (DOT), arrays of optical sources and detectors are used to perform functional neuroimaging experiments on task-based or resting state modes of human brain function.
Analyses are broken into several pipelines:
  B: modeling of the tissue shape, optical property distribution, and source/detector locations
  C: modeling of the light emission, diffusion, and detection through the head
  D: preprocessing of the raw source-detector measurements
  E: reconstruction and spectroscopy of the preprocessed data and light model into a functional neuroimaging volume
  F: post-processing analysis of these results.

## Slide 5: Generating a Light Model
This tutorial will follow the Basic_Light_Modeling_with_Hummingbird_Tutorial_220714.m Matlab file which can be found in the Documentation/Scripts folder.
The tutorial will:
  Generate a mesh from a segmented volume
  Place an array of sources and detectors on the mesh
  Set multiple parameters for the model
  Calculate the sensitivity A for this model
We will use a segmented volume, Segmented_MNI152nl_on_MNI111, found in Support_Files/Atlases, to generate the head model.
We will use the visual pad, Pad_AdultV24x28.mat, found in Support_Files/Pad.

## Slide 6
NeuroDOT Flowchart for Light Model Generation
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
These steps are repeated for a low-density and high-density mesh

## Slide 7
Running the script – Author’s Note
How the author runs the script
  Highlight and execute each chunk of code
    Ex: highlight lines 64-69  right click  hit “Evaluate Selection”
Alternate way of running the script
  For each section (beige) hit ctrl+enter
  This isn’t how the author does it, but it will still work
Note: DO NOT HIT THE GREEN “RUN” BUTTON to run this script
MATLAB versions to use
  Optimal version: MATLAB 2018
    The version that hummingbird was developed in
    Recommended to not use versions prior to 2018
  MATLAB 2020/2020b
    Hummingbird will run VERY SLOW when using newer versions of MATLAB, as AppDesigner was updated in 2020 and Hummingbird was developed prior to 2020
Example
section

## Slide 8
NeuroDOT Set-Up – In Script
The first section of code in this script will add the required toolboxes to your path and allow you to navigate to your desired output directory from within the script
Some parameters for visualizations made in the script are set up here as well
%Set Output Directory
%Change the path on the following line to whatever output directory you want to use
outputdir = ‘’; % put path to directory in between quotes
cd(outputdir)
%Parameter Initialization
p.Cmap='jet'; p.Scale=5; p.Th.P=0; p.Th.N=-p.Th.P; p.PD=1; p.BG=[0,0,0];
pM.orientation='coord'; pM.Cmap.P='gray'; pM.EdgeColor='none';
pS = pM; pS = rmfield(pS, 'orientation'); % make copy of params, to be used with PlotSlices

## Slide 9
Load a Segmented volume and visualize
[mask,infoT1]=LoadVolumetricData(['Segmented_MNI152nl_on_MNI111'],[],'4dfp');
p.Cmap='jet';p.Scale=5;p.Th.P=0;p.Th.N=-p.Th.P;p.PD=1;p.BG=[0,0,0];
PlotSlices(mask,infoT1,p) % Visualize the segmented mask
Note that PlotSlices.m generates an interactive visualization of the volume. Simply press ‘q’ or the middle mouse button to release the figure.
MRI Acquisition
Image Segmentation
5 – Scalp
4 – Skull
3 – Gray Matter
2 – White Matter
1 – CSF
0 - Background

## Slide 10
Generate a low-density mesh and visualize
The low-density mesh will be used to efficiently align the optical array to its position on the head.
% Parameters for generating your mesh
meshname=[‘LD_Mesh']; % Provide a name for your mesh name here
param.facet_distance=5; % Node position error tolerance at boundary
param.facet_size=3; % boundary element size parameter
param.cell_size=5; % Volume element size parameter
param.info=infoT1;
param.Offset=[0,0,0];
param.r0=5; % nodes outside of mask must be set to scalp==5;
param.CheckMeshQuality=0;
param.Mode=1; % make simple mesh with no region labels
tic;meshLD=NirfastMesh_Region(mask,meshname,param);toc
PlotMeshSurface(meshLD,pM) %visualize in index (non-negative) space
view([135,30])
meshLD.nodes=change_space_coords(meshLD.nodes,infoT1,'coord');
PlotMeshSurface(meshLD,pM) %visualize in coordinate space
view([135,30])
% Look at the inside
m1.nodes=meshLD.nodes;m1.elements=meshLD.elements;
m2=CutMesh(m1,find(m1.nodes(:,2)>0)); %cut LD mesh in half
[Ia,Ib]=ismember(m2.nodes,mesh.nodes,'rows');Ib(Ib==0)=[];
PlotMeshSurface(m2,pM); %visualize back half of LD mesh
view([135,30])
MRI Acquisition
Image Segmentation
LD Mesh Generation
“index space”
is non-negative
“coordinate space”
Usually has zeros in center of volume
A view inside

## Slide 11
Place the optodes on the LD mesh
Load in a Pad file for a grid design of interest, here we use the AdultV24x28 array.
padname='AdultV24x28';
load(['Pad_',padname,'.mat']); % Pad file
tpos=cat(1,info.optodes.spos3,info.optodes.dpos3);
Ns=size(info.optodes.spos3,1);
Nd=size(info.optodes.dpos3,1);
rad=info;
paramsFoci.color=cat(1,repmat([1,0,0],Ns,1),repmat([0,0,1],Nd,1));
paramsFoci.color(1,:) = [1 0.4 0.6]; %pink for s1
paramsFoci.color(Ns+1,:) = [0.3010, 0.7450, 0.9330]; %light blue for d1
paramsFoci.radius = 2; %set radius of spheres to 2
PlotSD(tpos(1:Ns,:),tpos((Ns+1):end,:),'norm');
PlotCap(info)
PlotMeshSurface(meshLD,pM);PlotSD(tpos(1:Ns,:),tpos((Ns+1):end,:),'norm',gcf);
view([135,30])
Note: This Pad file is a starting point and may not align with the mesh right away.
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement

## Slide 12
Place the optodes on the LD mesh
Adjust array position via the Hummingbird App by altering parameters and re-visualizing mesh and array:
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
First, reset the grid position
  Select: “to Behind the Head” from the drop down
  The view will then update
Then, choose a view that will best show the location of the array relative to the mesh
  XY view is great for this
Input new values for parameters and Update Figure
  dx, dy, dz: translate the array
  dxTh, dyTh, dzTh: rotate the array
  Scale: size the array
  flipX, flipY, flipZ: flip the array
Repeat until array is ~5mm off the head for the whole pad but is lined up with the mesh. After following the upcoming tutorial, the mesh and grid should look like the image on the right.
Note: The following 15 slides are an in-depth tutorial on how to use the Hummingbird app

## Slide 13: Hummingbird Tutorial 1
Run the following code to create the DataStorage class and open the Hummingbird app
ds = DataStorage();
% input structure
ds.dI.tpos = tpos;
ds.dI.mesh = meshLD;
ds.dI.pM = pM;
ds.dI.paramsFoci = paramsFoci;
ds.dI.Ns = Ns;
ds.dI.rad = rad;
myapp = Hummingbird(ds);
while isvalid(myapp); pause(0.1); end
Default view on Hummingbird launch

## Slide 14: Hummingbird Tutorial 2
Change the view to locate the pad
  The pad is the grid comprised of red and blue spheres
  Sometimes the pad may be sitting inside of the mesh upon startup
Use XY view to start

## Slide 15: Hummingbird Tutorial 3
Reset the pad to behind the head to visualize it
  This is a visual pad, and covers the occipital lobe and sits on the posterior surface of the mesh
Notice that the values changed for dx, dy, and dz
  These are the changes in x, y, and z from the original position of the pad when Hummingbird was initialized
  Let’s reset the view and begin to line up the mesh

## Slide 16: Hummingbird Tutorial 4
Hit “XZ View” to return to the initial view, which covers the back of the mesh
Now, lets rotate the pad so we can line it up

## Slide 17: Hummingbird Tutorial 5
Rotate the pad by modifying, dxTh, dyTh, or dzTh with the angle IN DEGREES that you want to rotate the pad by
  Here we rotated the pad around the x axis by 90 degrees, dxTh = -90
  Hit “Update Figure” to update the grid’s orientation
The blue and pink sources and detectors on the pad correspond to S/D 1, you can check this with the pad layout visualization generated earlier in the script
Additionally, the “View SD Labels” button will switch the pad from spheres to SD numbers which makes verifying against the layout visualization easier

## Slide 18: Hummingbird Tutorial 6
After hitting the “View SD Labels” button, the grid will show up as numbers instead of spheres
  Also, the button will change to say “View SD Spheres”
The blue and pink sources and detectors on the pad correspond to S/D 1, you can check this with the pad layout visualization generated earlier in the script

## Slide 19: Hummingbird Tutorial 9
Adjust the “Scale” to change the distance between optodes
Recommended Scale = 1.1
Hit “Update Figure” after setting the scale

## Slide 20: Hummingbird Tutorial 7
The bottom row of optodes should line up with the very top of the ear.
Adjust “dz” to do this
  Changed from -70.32 to -60.32
Hit “Update Figure” after adjusting dz

## Slide 21: Hummingbird Tutorial 8
Change to “YZ view”
Visualization check
  The pad should be no more than 10mm away from the mesh
  Let’s move it closer

## Slide 22: Hummingbird Tutorial 10
Visualization check
  The pad should be no more than 10mm away from the mesh
  Let’s move it closer
dy changed from -147.9 to -120.9
Remember to hit “Update Figure” after changing dy
Let’s visualize in XY view one more time to check the distance

## Slide 23: Hummingbird Tutorial 11
One last visualization check
  Hit “XY View”
We can move the pad slightly closer. Let’s adjust “dy” one last time

## Slide 24: Hummingbird Tutorial 12
Dy changed from -120.9 to -115.9
  Remember to hit “Update Figure” afterwards
Looks good! Now time to relax the grid onto the mesh
Hit the “Relax Mesh” button

## Slide 25: Hummingbird Tutorial 13
After hitting “Relax Mesh”, confirm that you want to use the current scale parameter
Current scale parameter = 1.1
After hitting “Done”, the below two figures will pop up and start to update
Hummingbird is now relaxing the optodes onto the mesh

## Slide 26: Hummingbird Tutorial 14
Once optodes are relaxed onto the mesh, a dialogue box will show up on the Hummingbird app
Hit “OK” to close it

## Slide 27: Hummingbird Tutorial 15
To save the new optode locations to your MALTAB workspace, hit the “Finalize” button
This will also close Hummingbird
Note: on some versions of MATLAB, the finalize button will not be visible
  Simply close hummingbird and resume the script at the next line, the relaxed optodes will be saved to the DataStorage object and you will be able to access them
Congrats!! You’ve finished the Hummingbird tutorial!

## Slide 28
Generate smaller mask based on optode locations
Here, we get our relaxed grid position and visualize it on the LD mesh. Then, we scoop out the part of the mesh that the grid is sitting on. This scooped section will is the basis for the high-density mesh that we are about to create
% get relaxed optode positions from Hummingbird
tposNew = ds.dO.tpos2_relaxed;
% visualize mesh with relaxed optodes
pM.reg=0;
PlotMeshSurface(meshLD,pM);Draw_Foci_191203(tposNew,paramsFoci)
view(0,0)
%% Generate smaller mask based on optode locations
maskCrop=IceCreamScoopCap4Mask(infoT1,tposNew,mask,40); %40 = mm of mesh to scoop out
PlotSlices(maskCrop,infoT1,pS) % visualize smaller mask
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation

## Slide 29
Generate a high-density mesh and visualize
Take the scooped section of the LD mesh and turn it into a HD mesh
% If you get an error when running NirfastMesh_Region,
% try changing the mesh name and then re-run NirfastMesh_Region
meshname=['_HD_Mesh1']; % Provide a name for your mesh name here
param.facet_distance=2.0; % Node position error tolerance at boundary
param.facet_size=0.8; % boundary element size parameter
param.cell_size=1.0; % Volume element size parameter
param.Mode=0; % Set Mode=0 to make nirfast compliant mesh
tic;meshHD=NirfastMesh_Region(maskCrop,meshname,param);toc
%make copy of mesh that only contains nodes and elements for visualization purposes
visMeshHD.nodes = meshHD.nodes;
visMeshHD.elements = meshHD.elements;
PlotMeshSurface(visMeshHD,pM);view([70,60]) %Visualize in index space
% Put coordinates back in coordinate space meshHD.nodes=change_space_coords(meshHD.nodes,infoT1,'coord'); %for actual mesh visMeshHD.nodes=change_space_coords(visMeshHD.nodes,infoT1,'coord'); %for mesh that’s visualized PlotMeshSurface(visMeshHD,pM);view([70,60]) %Visualize in coordinate space
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
“index space”
is non-negative
“coordinate space”
Usually has zeros in center of volume

## Slide 30
Generate a high-density mesh and visualize
Notes and possible errors
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Notes on HD mesh generation
Parameters
  Facet_distance = error tolerance at boundary for node positions
    Typically set to 2.0
  Facet_size = size of surface triangles of mesh
    Want to keep this as small as possible bc we want the surface to be highly faceted
    Typically set to 0.8
  Cell_size = size of internal tetrahedra
    This can be changed depending on if the total number of nodes in the mesh is > 999,999
    Typically set to 1.0
Possible Error
File I/O Error:
  Solution: delete all mesh files in your output directory, change the meshname variable to a new name and re-run
    Ex: ‘_HD_Mesh1’  ‘_HD_Mesh0’
  Make sure you delete these files from the file manager, and NOT within matlab
  If the above doesn’t work, create an entirely new output directory, make that directory your current folder, then re-run lines 149-168
Files to delete if File I/O Error occurs
Pop-up box for File I/O Error

## Slide 31
Place the optodes on the HD mesh
% get relaxed optode positions from Hummingbird
% input structure
ds_HD = DataStorage();
ds_HD.dI.tpos = tposNew;
ds_HD.dI.mesh = visMeshHD;
ds_HD.dI.pM = pM;
ds_HD.dI.paramsFoci = paramsFoci;
ds_HD.dI.Ns = Ns;
ds_HD.dI.rad = rad;
% run hummingbird
myapp=Hummingbird(ds_HD);
while isvalid(myapp);pause(0.1);end
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Optode Placement
This time, since the HD mesh is in the same space that the LD mesh was, the grid will already be lined up and sitting on the mesh when hummingbird starts. You’ll need to move it away from the mesh and re-relax the grid back onto the mesh
Change the view to “XY View”
Move the grid off the mesh by setting dy = -10
Adjust the scale to 1.1 (this is the scale we used for the LD mesh too)
Hit “Update Figure”
  Hummingbird’s GUI should now look like the image below
Hit “Relax Mesh” and “Finalize” after the grid has been re-relaxed onto the mesh

## Slide 32
Update pad info structure after running Hummingbird
MRI Acquisition
Image Segmentation
Mesh Generation
Optode Placement
After the pad has been relaxed onto the HD mesh, we need to update the info structure with the new optode locations
info.optodes.spos3=tposNew_HD(1:Ns,:);
info.optodes.dpos3=tposNew_HD((Ns+1):end,:);
m=0;
for d=1:Nd
for s=1:Ns
m=m+1;
info.pairs.r3d(m)=norm(tposNew_HD(s,:)-tposNew_HD(d+Ns,:));
info.pairs.r3d(m+(Ns*Nd))=info.pairs.r3d(m);
end
end
% histogram of number of measurements per SD vs radius between SD pairs figure;histogram(info.pairs.r3d,1000);xlabel('R_S_D');ylabel('N_m_e_a_s');
% save pad file
gridname=padname;
save(['Pad_',meshname,'_',gridname,'_', '.mat'],'info')
For both LD and HD meshes

## Slide 33
Calculate the sensitivity (1 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Prepare mesh and optode information for NIRFAST, then visualize
gridname=padname;
mesh=PrepareMeshForNIRFAST(mesh,[meshname,'_',gridname],tposNew);
PlotMeshSurface(mesh,pM);PlotSD(mesh.source.coord(1:Ns,:),...
mesh.source.coord((Ns+1):end,:),'render',gcf);
Visualize cap and mesh with portion cut out to make extra sure that optodes are not placed too far within the mesh
m3=CutMesh(mesh,intersect(find(mesh.nodes(:,3)>0),find(mesh.nodes(:,1)>0)));
[Ia,Ib]=ismember(m3.nodes,mesh.nodes,'rows');Ib(Ib==0)=[];
m3.region=mesh.region(Ib);
PlotMeshSurface(m3,pM);PlotSD(mesh.source.coord(1:Ns,:),...
mesh.source.coord((1+Ns):end,:),'render',gcf);
For both LD and HD meshes
1.
2.

## Slide 34
Calculate the sensitivity (2 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
3. Set parameters for light modeling
flags.tag=[padname,'_on_',meshname,'_test'];
flags.gridname=gridname;
flags.meshname=meshname;
flags.head='info';
flags.info=infoT1; % Your T1 info file
flags.gthresh=1e-5; % Voxelation threshold in G
flags.voxmm=2; % Voxelation resolution (mm)
flags.labels.r1='csf'; % Regions for optical properties
flags.labels.r2='white';
flags.labels.r3='gray';
flags.labels.r4='bone';
flags.labels.r5='skin';
flags.op.lambda=[750,850]; % Wavelengths (nm)
flags.op.mua_skin=[0.0170,0.0190]; % Baseline absorption
flags.op.mua_bone=[0.0116,0.0139];
flags.op.mua_csf=[0.0040,0.0040];
flags.op.mua_gray=[0.0180,0.0192];
flags.op.mua_white=[0.0167,0.0208];
flags.op.musp_skin=[0.74,0.64]; % Baseline reduced scattering coeff
flags.op.musp_bone=[0.94,0.84];
flags.op.musp_csf=[0.3,0.3];
flags.op.musp_gray=[0.8359,0.6726];
flags.op.musp_white=[1.1908,1.0107];
flags.op.n_skin=[1.4,1.4]; % Index of refration
flags.op.n_bone=[1.4,1.4];
flags.op.n_csf=[1.4,1.4];
flags.op.n_gray=[1.4,1.4];
flags.op.n_white=[1.4,1.4];
flags.srcnum=Ns; % Number of sources
flags.t4=eye(4); % T1/dim to MNI atlas *** change this to register your vol to atlas
flags.t4_target='MNI'; % string
flags.makeA=1; % don't make A, just make G
flags.Hz=0;
if flags.Hz, flags.tag = [flags.tag,'FD']; end
For both LD and HD meshes

## Slide 35
Calculate the sensitivity (3 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
4. Calculate Green’s functions, voxelate, create sensitivity matrix, adjust bookkeeping
Ti=tic;[A,dim,Gsd]=makeAnirfast(mesh,flags); % size(A)= [Nwl, Nmeas, Nvox]
disp(['<makeAnirfast took ',num2str(toc(Ti))])
%Gsd vs Rsd provides a simulated light fall-off figure;semilogy(info.pairs.r3d(1:(Ns*Nd)),Gsd,'*');
legend({num2str(flags.op.lambda')})
xlabel('R_S_D [mm]');ylabel('G_S_D');xlim([0,100])
If makeAnirfast is successful, info like the below screenshot should show up in your command window
For both LD and HD meshes

## Slide 36
Package Sensitivity Matrix with associated info
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
5. Arrange A to have the same measurement list structure as the data:
[Nwl,Nmeas,Nvox]=size(A);
A=reshape(permute(A,[2,1,3]),Nwl*Nmeas,Nvox);
6. Place spatial information about light model in info.tissue structure:
info.tissue.dim=dim;
info.tissue.affine=flags.t4;
info.tissue.infoT1=infoT1;
info.tissue.affine_target='MNI';
info.tissue.flags=flags;
temp_pairs=struct;
temp_pairs.Src=info.pairs.Src;
temp_pairs.Det=info.pairs.Det;
temp_pairs.NN=info.pairs.NN;
temp_pairs.WL=info.pairs.WL;
temp_pairs.lambda=info.pairs.lambda;
temp_pairs.Mod=info.pairs.Mod;
temp_pairs.r2d=info.pairs.r2d;
temp_pairs.r3d=info.pairs.r3d;
info.pairs=temp_pairs;
save(['A_',flags.tag,'.mat'],'A','info','-v7.3’)
For both LD and HD meshes

## Slide 37
Package Sensitivity Matrix with associated info
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
6. Package and save A where all source-detector separations are less than 6cm
For high-density pads, only measurements from source-detector (SD) pairs that have a separation distance within 6cm are generally used. These overlapping measurements are what provide the image quality of the HD arrays. The Gsd vs Rsd scatterplot is a helpful tool for visualizing what measurements meet the 6cm cutoff. A red box has been placed around the measurements within the 6cm cutoff.
% Get A with all SD separations within 6cm
keep=info.pairs.r3d<=60;
temp=struct;
temp.Src=info.pairs.Src(keep);
temp.Det=info.pairs.Det(keep);
temp.NN=info.pairs.NN(keep);
temp.WL=info.pairs.WL(keep);
temp.lambda=info.pairs.lambda(keep);
temp.Mod=info.pairs.Mod(keep);
temp.r2d=info.pairs.r2d(keep);
temp.r3d=info.pairs.r3d(keep);
info.pairs=temp;
A=A(keep,:); %A with all SD within 6cm
save(['A_r3d_lth_60_',flags.tag,'.mat'],'A','info','-v7.3') %save A with all SD within 6cm
For both LD and HD meshes
Gsd vs Rsd

## Slide 38
Validate the sensitivity (1 of 2)
Visualize an example measurement’s sensitivity using the A matrix containing measurements from SD’s with separations < 6cm. We’ll start with a single source detector pair, and visualize the full sensitivity profile (FFR) on the next slide.
% Visualize aspects of sensitivity profile
t1=affine3d_img(mask,infoT1,dim,eye(4)); % put anatomical volume in dim space
keep=info.pairs.WL==2 & info.pairs.Src==1 & info.pairs.Det==6;
foo=squeeze(A(keep,:)); % Single meas pair
fooV=Good_Vox2vol(foo',dim);
fooV=fooV./max(fooV(:));
fooV=log10(1e2.*fooV); % top 2 o.o.m.
pA.PD=1;pA.Scale=2;pA.Th.P=0;pA.Th.N=-pA.Th.P;
PlotSlices(t1,dim,pA,fooV)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes
Single SD pair: source 1 detector 6

## Slide 39
Validate the sensitivity (2 of 2)
Visualize the flat field reconstruction
keep=(info.pairs.WL==2 & info.pairs.r2d<=40);
a=squeeze(A(keep,:));
iA=Tikhonov_invert_Amat(a,0.01,0.1);
iA=smooth_Amat(iA,dim,5); % use smoothing parameter of 5 here
ffr=makeFlatFieldRecon(a,iA);
fooV=Good_Vox2vol(ffr,dim);
fooV=fooV./max(fooV(:));
pA.PD=1;pA.Scale=1;pA.Th.P=1e-2;pA.Th.N=-pA.Th.P;
PlotSlices(t1,dim,pA,fooV)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes

## Slide 40
Put it all together
As a final check, let’s visualize the alignment of the LD mesh, HD mesh, optode array, and cortical meshes all together
% View alignment
Cortical_mesh = load(['MNI164k_big.mat']);
Anat.CtxL = Cortical_mesh.MNIl;
Anat.CtxR = Cortical_mesh.MNIr;
pA0l=struct;
figure('Color','k','Position',[500,100,1050,1000])
pM1.fig_handle=gca;pM1.FaceColor=[1,1,1].*0.5;
pM1.SpecularStrength=0.25;pM1.DiffuseStrength=0.5;
PlotMeshSurface(meshHD,pM1) % HD mesh
pA0l.fig_handle=gca;
pA0l.FaceColor=[1,1,1].*0.5;pA0l.EdgeColor='none';
pA0l.AmbientStrength=0.25;pA0l.DiffuseStrength=0.25;
pA0l.SpecularStrength=0.025;
PlotMeshSurface(Anat.CtxL,pA0l) % Cortical Surfaces
PlotMeshSurface(Anat.CtxR,pA0l)
pA2.fig_handle=gca;
pA2.FaceAlpha=0;
pA2.EdgeAlpha=0.5;
pA2.EdgeColor=[1,1,1].*0.25;
PlotMeshSurface(meshLD,pA2) % LD mesh
set(gca,'Color','k')
view([90,0]) % display in lateral view
Draw_Foci_191203(cat(1,info.optodes.spos3,info.optodes.dpos3),paramsFoci) % Array
axis off
% For posterior view uncomment the following line
% view([0,0])

## Slide 41: Appendix: Generating info for Pad file (1 of 3)
The Pad file for a given grid design contains a structure variable info that contains the spatial and topological information for the array as follows:
  info.optodes
    Structure containing spatial coordinates for sources and detectors in array
    info.optodes.spos2 – 2D coordinates of each sources
    info.optodes.spos3 – 3D coordinates of each source
    info.optodes.dpos2 – 2D coordinates of each detector
    info.optodes.dpos3 – 3D coordinates of each detector
  info.pairs
    Table containing topological information for each source-detector measurement pair
    Each row of this table corresponds to a measurement and must be in the same order as the optical data for valid processing
    Generally, full list of sources and detectors is repeated for each wavelength/lambda.
    info.pairs.Src – source in measurement
    info.pairs.Det – detector in measurement
    info.pairs.NN – Nearest Neighbor class in array (e.g., info.pairs.NN=1 means this measurement is one of the closest in the grid design)
    info.pairs.WL – wavelength index of measurement
    info.pairs.lambda – actual wavelength of measurement (e.g., info.pairs.lambda-=750 is a 750 nm measurement)
    info.pairs.Mod – Modulation type of measurement (e.g., ‘CW’, or, ‘FD’)
    info.pairs.r2d – 2D source-detector distances
    info.pairs.r3d – 3D source-detector distances

## Slide 42: Appendix: Generating info for Pad file (2 of 3)
If you already have a grid design with 2D and 3D source and detector locations, saved as the following:
  grid.spos2
  grid.spos3
  grid.dpos2
  grid.dpos3
To generate the topology structure for your grid:
  dr=5; % Distance separating 'nearest neighbor' groups
  Rad=Grid2Radius_180824(grid,dr);
Once you have the grid and radius set up (as above), the info structure can be generated with the following:
  grid.name='Your_Grid_Name';
  params.lambda=[750,850]; % Wavelengths for your system
  params.Mod='CW'; % Modulation strategy
  [info]=Generate_Info_from_Grid_Rad(grid,Rad,params);
Then just save the Pad file for future use:
  save(['Pad_',info.optodes.CapName],'info')

## Slide 43: Appendix: Generating info for Pad file (3 of 3)
If instead you are designing a grid from scratch in style of our visual cap:
  Ndy=4; % Num detectors in y-direction
  Ndx=7; % Num detectors in x-direction
  NN1sep=13; % Closest source-detector Separation
  Srad=80; % Radius of spherical head model
  [grid,Rad]=Make_Sphere_Cap(Ndy,Ndx,NN1sep,Srad);
As described above,
  grid contains information on the positions of the sources and detectors.
  Rad contains information on the topology of the array.
Visualize the grid in 2D and 3D
  PlotSD(grid.spos2,grid.dpos2,'norm'); % Visualize 2D positions
  PlotSD(grid.spos3,grid.dpos3,'norm'); % Visualize 3D positions
Then set parameters, generate info file as on previous slide, and save Pad file for future use.

## Slide 44: That’s It (For Now)
Congratulations! You have finished the NeuroDOT Light Modeling Pipeline Tutorial for the basic 24x28 visual pad using Hummingbird.
For further questions or more information, please consult the NeuroDOT User Manual and the various Appendices.
NeuroDOT Team:
  Adam Eggebrecht (aeggebre@wustl.edu)
  Emma Speh (espeh@wustl.edu)
  Ari Segel (ari@wustl.edu)