## Slide 1: NeuroDOT
Tutorial:
Generating a Light Model

## Slide 2
NeuroDOT Set-Up - Checklist
Before starting any tutorial, please make sure you’ve done the following
Download the appropriate versions of matlab, neuroDOT, and NIRFAST
  Matlab: 2020b
  NeuroDOT toolbox from NITRC
    https://www.nitrc.org/projects/neurodot/
  NIRFASTer-master
    https://github.com/nirfaster/NIRFASTer
Set your path so that all required folders and subfolders are included
  Folders w/subfolders to include
    NeuroDOT toolbox
    NIRFASTer-master
    Your personal directory
  Folders to include
    Output directory

## Slide 3
NeuroDOT Set-Up – Output Directory
Once you’ve set your path, the next thing you should do is set your output directory.
  The output directory must be located within your path
  If it isn’t already in your path, you must add it to your path
You can set your own path to an output directory in the line shown above
An output directory specifies the location where files will be saved
  MATLAB, by default, will save files to the current directory
The following line of code will change your current directory to the output directory that you just specified 
Now you can save outputs directly to your output directory
Example output directory

## Slide 4: Diffuse Optical Tomography
In diffuse optical tomography (DOT), arrays of optical sources and detectors are used to perform functional neuroimaging experiments on task-based or resting state modes of human brain function.
Analyses are broken into several pipelines:
  B: modeling of the tissue shape, optical property distribution, and source/detector locations
  C: modeling of the light emission, diffusion, and detection through the head
  D: preprocessing of the raw source-detector measurements
  E: reconstruction and spectroscopy of the preprocessed data and light model into a functional neuroimaging volume
  F: post-processing analysis of these results.

## Slide 5: Generating a Light Model
This tutorial will follow the Basic_Light_Modeling_with_AlignMe_Tutorial_220928.m Matlab file which can be found in the Documentation/Scripts folder.
The tutorial will:
  Generate a mesh from a segmented volume
  Place an array of sources and detectors on the mesh
  Set multiple parameters for the model
  Calculate the sensitivity A for this model
We will use a segmented volume, Segmented_MNI152nl_on_MNI111_nifti, found in Support_Files/Atlases, to generate the head model.
We will use the visual pad, Pad_AdultV24x28.mat, found in Support_Files/Pad.

## Slide 6
NeuroDOT Flowchart for Light Model Generation
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
These steps are repeated for a low-density and high-density mesh

## Slide 7
Running the script – Author’s Note
How the author runs the script
  Highlight and execute each chunk of code
    Ex: highlight lines 64-69  right click  hit “Evaluate Selection”
Alternate way of running the script
  For each section (beige) hit ctrl+enter
  This isn’t how the author does it, but it will still work
Note: DO NOT HIT THE GREEN “RUN” BUTTON to run this script
MATLAB versions to use
  Optimal version: MATLAB 2020b
    The version that AlignMe was developed in
NIRFAST version to use
  NIRFASTer, the newest release of NIRFAST
  Found here: https://github.com/nirfaster/NIRFASTer
Example
section

## Slide 8
NeuroDOT Set-Up – In Script
The first section of code in this script will have you set your output directory and allow you to navigate to your desired output directory from within the script
Some parameters for visualizations made in the script are set up here as well
%Set Output Directory
%Change the path on the following line to whatever output directory you want to use
outputdir = ‘’; % put path to directory in between quotes
cd(outputdir)
%Parameter Initialization
p.Cmap='jet'; p.Scale=5; p.Th.P=0; p.Th.N=-p.Th.P; p.PD=1; p.BG=[0,0,0];
pM.orientation='coord'; pM.Cmap.P='gray'; pM.EdgeColor='none';
pS = pM; pS = rmfield(pS, 'orientation'); % make copy of params, to be used with PlotSlices
23.
24.
25.
26.
28.
29.
30.
31.

## Slide 9
Load a Segmented volume and visualize
[mask,infoT1]=LoadVolumetricData(['Segmented_MNI152nl_on_MNI111_nifti'],[],’nii');
PlotSlices(mask,infoT1,p) % Visualize the segmented mask
Note that PlotSlices.m generates an interactive visualization of the volume. Simply press ‘q’ or the middle mouse button to release the figure.
MRI Acquisition
Image Segmentation
5 – Scalp
4 – Skull
3 – Gray Matter
2 – White Matter
1 – CSF
0 - Background
35.
36.

## Slide 10
The low-density mesh will be used to efficiently align the optical array to its position on the head
% Parameters for generating your mesh
meshname=[‘LD_Mesh']; % Provide a name for your mesh name here
param.facet_distance=5; % Node position error tolerance at boundary
param.facet_size=3; % boundary element size parameter
param.cell_size=5; % Volume element size parameter
param.info=infoT1;
param.Offset=[0,0,0];
param.r0=5;
param.CheckMeshQuality=0;
param.Mode=1;
tic;meshLD=NirfastMesh_Region(mask,meshname,param);toc
Parameters of note
  Facet_distance = error tolerance at boundary for node positions
    Typically set to 2.0
  Facet_size = size of surface triangles of mesh
    Want to keep this as small as possible bc we want the surface to be highly faceted
    Typically set to 0.8
  Cell_size = size of internal tetrahedra
    This can be changed depending on if the total number of nodes in the mesh is > 999,999
    Typically set to 1.0
Generate a low-density mesh and visualize
MRI Acquisition
Image Segmentation
LD Mesh Generation
“index space”
is non-negative
“coordinate space”
Usually has zeros in center of volume
A view inside
43.
44.
45.
46.
47.
48.
49.
50.
51.
52.
Run lines 54-74 to generate visualizations

## Slide 11
Depending on your starting volume, the low-density mesh generation process will differ
Atlas based mesh (what this tutorial does)
  Set param.Offset = [0,0,0];
  Then use change_space_coordinates()
Subject specific based mesh
  Set param.Offset = InfoT1.center; (that’s it)
% Parameters for generating your mesh
meshname=[‘LD_Mesh']; % Provide a name for your mesh name here
param.facet_distance=5; % Node position error tolerance at boundary
param.facet_size=3; % boundary element size parameter
param.cell_size=5; % Volume element size parameter
param.info=infoT1;
param.Offset=[0,0,0];
param.r0=5;
param.CheckMeshQuality=0;
param.Mode=1;
tic;meshLD=NirfastMesh_Region(mask,meshname,param);toc
meshLD.nodes=change_space_coords(meshLD.nodes,infoT1,'coord’);
Atlas vs Subject Specific LD Mesh Generation
MRI Acquisition
Image Segmentation
LD Mesh Generation
“index space”
is non-negative
“coordinate space”
Usually has zeros in center of volume

## Slide 12
Make sure your mesh is in the same space as your volume
In PlotSlices, place crosshairs at the tip of the nose
  Note the X, Y, and Z coordinates for those slices
Generate LD mesh with your chosen parameters
  Param.Offset = [0,0,0]; or Param.Offset = infoT1;
  Use change_space_coordinates() or don’t use it
In atlas figure, click on the tip of the nose of the mesh
  Note the X, Y, and Z coordinates for those slices
Repeat the above 2 steps, with different parameters if the X, Y, Z coordinates are not similar
Validating LD Mesh Generation
MRI Acquisition
Image Segmentation
LD Mesh Generation
Atlas Volume
+
No CSC
Atlas Volume
+
CSC
Volume coords
X = -1
Y = 83
Z = -63
Coords are similar to volume!

## Slide 13
Load in a Pad file for a grid design of interest, here we use the AdultV24x28 array
padname='AdultV24x28';
load(['Pad_',padname,'.mat']); % Pad file
tpos=cat(1,info.optodes.spos3,info.optodes.dpos3);
Ns=size(info.optodes.spos3,1);
Nd=size(info.optodes.dpos3,1);
rad=info;
paramsFoci.color=cat(1,repmat([1,0,0],Ns,1),repmat([0,0,1],Nd,1));
paramsFoci.color(1,:) = [1 0.4 0.6]; %pink for s1
paramsFoci.color(Ns+1,:) = [0.3010, 0.7450, 0.9330]; %light blue for d1
paramsFoci.radius = 2; %set radius of spheres to 2
Note: This Pad file is a starting point and may not align with the mesh right away.
Place the optodes on the LD mesh
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
80.
81.
82.
83.
84.
85.
87.
88.
89.
90.
93.
94.
98.
99.
Draw_Foci_191203(tpos, paramsFoci);view([-40,30])
PlotCap(info)
PlotMeshSurface(meshLD,pM);Draw_Foci_191203(tpos, paramsFoci);
view([0,90])

## Slide 14
AlignMe Functionality
AlignMe allows for the user to scale the mesh to participant specific space by using
  Participant Head Circumference
  Participant Z line distance (nasion to inion)
These are contained within the Atlas Scaling box
Additionally, to facilitate easier alignment for participant specific situations, photos of the capfit can be utilized, and will show up where the Brain Light Lab logo is located on the GUI
  Rotate CW, Rotate CCW: rotates photo
  Previous, Next: scroll through available photos
Participant photos can be loaded in by setting the path to the directory containing photos. AlignMe will pull all files of the designated image type
  Path to photos: ds.dI.photopath
  Type of image: ds.dI.imageType
Note: The following 19 slides are an in-depth tutorial on how to use the AlignMe app
  The mesh will not be scaled in this tutorial
  There will not be any participant photos in this tutorial
You can find the following slides in a standalone tutorial titled: NeuroDOT_Tutorial_AlignMe
  Tutorial can be found in the Documentation/Scripts folder of the NeuroDOT toolbox

## Slide 15: AlignMe Tutorial 1
% Create an instance of our custom DataStorage HANDLE class
ds = DataStorage();
% Input structure
ds.dI.tpos = tpos;
ds.dI.mesh = meshLD;
ds.dI.pad = info;
ds.dI.pM = pM;
ds.dI.paramsFoci = paramsFoci;
ds.dI.Ns = Ns;
ds.dI.photoPath = ‘’;
ds.dI.imageType = '.jpeg';
ds.dI.atlasFiducials = atlasFiducials;
% Run AlignMe
myapp=AlignMe_2020b(ds);
while isvalid(myapp); pause(0.1); end
Default view on AlignMe launch
113.
114.
116.
117.
118.
119.
120.
121.
122.
123.
124.
125.
131.
132.

## Slide 16: AlignMe Tutorial 2
Change the view to locate the pad
  The pad is the grid comprised of red and blue spheres or cylinders
  Sometimes the pad may be sitting inside of the mesh upon startup
Use XY view to start

## Slide 17: AlignMe Tutorial 3
Reset the pad to behind the mesh to visualize it
  This is a visual pad which covers the occipital lobe and sits on the posterior surface of the mesh
Notice that the values changed for dx, dy, and dz
  These are the changes in x, y, and z from the original position of the pad when AlignMe was initialized
  Let’s reset the view and begin to line up the mesh

## Slide 18: AlignMe Tutorial 4
Hit “XZ View” to navigate to a view which covers the back of the mesh
Now, lets rotate the pad so we can line it up

## Slide 19: AlignMe Tutorial 5
Rotate the pad by modifying, x rot, y rot, or z rot, with the angle IN DEGREES that you want to rotate the pad by
  Here we rotate the pad around the x axis by 90 degrees, x rot = -90
To update the array’s position after changing any value in the Array Alignment box, such as x rot: simply hit enter, tab, or click somewhere else on the app
The blue and pink sources and detectors (S/D) on the pad correspond to S/D 1, you can check this with the pad layout visualization generated earlier in the script
  Line 89: PlotCap(info) %2D plot of pad

## Slide 20: AlignMe Tutorial 6
Instead of using cylinders to denote the optodes, there’s an option for using spheres. Simply, flip the switch
  Switch is located above the “Atlas Scaling” box
  Future versions of AlignMe will allow for the user to specify spheres or cylinders as an input in the DataStorage handle class
The blue and pink sources and detectors on the pad correspond to S/D 1, you can check this with the pad layout visualization generated earlier in the script

## Slide 21: AlignMe Tutorial 7
Additionally, to better view the optode numbers, click the “Turn Optode Labels On/Off” button
  If no numbers, button displays: Turn Optode Labels On
  If numbers, button displays: Turn Optode Labels Off
We’ll turn the optode labels off for the rest of the tutorial

## Slide 22: AlignMe Tutorial 8
Adjust the “Scale” to change the distance between optodes
Recommended Scale = 1.1
To update the array after changing any value in the Array Alignment box, such as Scale: simply hit enter, tab, or click somewhere else on the app

## Slide 23: AlignMe Tutorial 9
The bottom row of optodes should line up with the very top of the ear.
Adjust “dz” to do this
  Changed from -70.24 to -60.24
To update the array after changing any value in the Array Alignment box, such as dz: simply hit enter, tab, or click somewhere else on the app

## Slide 24: AlignMe Tutorial 10
Change to “YZ view”
Visualization check
  The pad should be no more than 10mm away from the mesh
  Let’s move it closer

## Slide 25: AlignMe Tutorial 11
Visualization check
  The pad should be no more than 10mm away from the mesh
  Let’s move it closer
dy changed from -147.8 to -120.8
  To update the array after changing any value in the Array Alignment box, such as dy: simply hit enter, tab, or click somewhere else on the app
Let’s visualize in XY view one more time to check the distance

## Slide 26: AlignMe Tutorial 12
One last visualization check
  Hit “XY View”
We can move the pad slightly closer. Let’s adjust “dy” one last time

## Slide 27: AlignMe Tutorial 13
Dy changed from -120.8 to -115.8
Looks good! Lets do one final sanity check of all 3 views
XZ view
XY view
YZ view

## Slide 28: AlignMe Tutorial 14
Looks good! Now time to relax the grid onto the mesh
Hit the “Relax Array” button
Then, the below two figures will pop up and start to update as the array is relaxed onto the surface of the mesh

## Slide 29: AlignMe Tutorial 15
Once optodes are relaxed onto the mesh, a dialogue box will show up on the AlignMe app
Hit “OK” to close the dialogue box

## Slide 30: AlignMe Tutorial 16
To save the new optode locations to your MALTAB workspace, hit the “Save Alignment” button

## Slide 31: AlignMe Tutorial Potential Errors
Sometimes, when clicking the “Save Alignment” button, the error shown on the right occurs
  Usually this is due to clicking the button multiple times before the app has finished running
To fix, just select 'Save Alignment' again
  Make sure you only click once
  Please be patient

## Slide 32: AlignMe Tutorial 17
A final dialogue box will open. Click OK to close AlignMe
You’ve finished the AlignMe tutorial!

## Slide 33: Examples with Common Commercial Array Designs
NIRX
Forehead pad
NIRX
Full head pad
Gowerlabs
Visual pad
Gowerlabs
Temporal pad

## Slide 34
Generate smaller mask based on optode locations
Here, we get our relaxed grid position and visualize it on the LD mesh. Then, we scoop out the part of the mesh that the grid is sitting on. This scooped section will is the basis for the high-density mesh that we are about to create
% get relaxed optode positions from AlignMe
tposNew = ds.dO.tpos2_relaxed;
% visualize mesh with relaxed optodes
pM.reg=0;
PlotMeshSurface(meshLD,pM);Draw_Foci_191203(tposNew,paramsFoci)
view(0,0)
% Get affine matrix that can be used to transform FROM participant space TO MNI space
affine_Subj2MNI = [ds.dO.affineTform, zeros(3,1)];
save('affine_matrix_Subject_to_MNI.mat', 'affine_Subj2MNI')
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
134.
135.
137.
138.
139.
140.
144.
145.
151.
152.
153.
%% Generate smaller mask based on optode locations
maskCrop=IceCreamScoopCap4Mask(infoT1,tposNew,mask,40); %40 = mm of mesh to scoop out
PlotSlices(maskCrop,infoT1,pS) % visualize smaller mask

## Slide 35
Generate a high-density mesh and visualize
Take the scooped section of the LD mesh and turn it into a HD mesh
% If you get an error when running NirfastMesh_Region,
% try changing the mesh name and then re-run NirfastMesh_Region
meshname=['_HD_Mesh1']; % Provide a name for your mesh name here
param.facet_distance=2.0; % Node position error tolerance at boundary
param.facet_size=0.8; % boundary element size parameter
param.cell_size=1.0; % Volume element size parameter
param.Mode=0; % Set Mode=0 to make nirfast compliant mesh
param.info=infoT1; % Make sure info is infoT1 like with LD mesh
param.Offset=[0,0,0];
tic;meshHD=NirfastMesh_Region(maskCrop,meshname,param);toc
%make copy of mesh that only contains nodes and elements for visualization purposes
visMeshHD.nodes = meshHD.nodes;
visMeshHD.elements = meshHD.elements;
PlotMeshSurface(visMeshHD,pM);view([70,60]) %Visualize in index space
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
“index space”
is non-negative
“coordinate space”
Usually has zeros in center of volume
157.
158.
159.
160.
161.
162.
163.
164.
165.
166.
168.
169.
170.
171.
175.
176.
177.
178.
% Put coordinates back in coordinate space
meshHD.nodes=change_space_coords(meshHD.nodes,infoT1,'coord'); %for actual mesh
visMeshHD.nodes=change_space_coords(visMeshHD.nodes,infoT1,'coord');%for mesh that’s visualized
PlotMeshSurface(visMeshHD,pM);view([70,60]) %Visualize in coordinate space
Notes: NOTE: WE ARE NOT HAVING THEM RUN THESE SLIDES

## Slide 36
Generate a high-density mesh and visualize
Notes and possible errors
Files to delete if File I/O Error occurs
Pop-up box for File I/O Error
Notes on HD mesh generation
Possible Errors
File I/O Error:
  Solution: delete all mesh files in your output directory, change the meshname variable to a new name and re-run
  Ex: ‘_HD_Mesh1’  ‘_HD_Mesh0’
  Make sure you delete these files from the file manager, and NOT within matlab
  If the above doesn’t work, create an entirely new output directory, make that directory your current folder, then re-run lines 149-168
Poor mesh quality
  Occurs if number of nodes in mesh is greater than 1,000,000
  This is why we make smaller HD meshes and tune params.cell size
Parameters
Cell_size = size of internal tetrahedra
  This can be changed depending on if the total number of nodes in the mesh is > 999,999
  Typically set to 1.0 for visual pad, and 2.0 for larger volume meshes i.e. full head
Facet_distance = error tolerance at boundary for node positions
  Typically set to 2.0 for HD meshes
Facet_size = size of surface triangles of mesh
  Want to keep this as small as possible, as we want the surface to be highly faceted
  Typically set to 0.8 for HD meshes

## Slide 37
Place the optodes on the HD mesh
%Create an instance of DataStorage class
ds_HD = DataStorage();
% Input structure
ds_HD.dI.tpos = tposNew;
ds_HD.dI.mesh = meshHD;
ds_HD.dI.pad = info;
ds_HD.dI.pM = pM;
ds_HD.dI.paramsFoci = paramsFoci;
ds_HD.dI.Ns = Ns;
ds_HD.dI.photoPath = '';
ds_HD.dI.imageType = '.jpeg'; ds_HD.dI.atlasFiducials = atlasFiducials;
% run AlignMe
myapp = AlignMe_2020b(ds_HD);
while isvalid(myapp);pause(0.1); end
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Optode Placement
This time, since the HD mesh is in the same space that the LD mesh was, the grid will already be lined up and sitting on the mesh when AlignMe starts. You’ll need to move it away from the mesh and re-relax the grid back onto the mesh
Change the view to “XY View”
Switch from cylinders to spheres
Adjust the scale to 1.1 (this is the scale we used for the LD mesh too)
Move the grid off the mesh by setting dy = -10
  The AlignMe GUI should now look like the image below
Hit “Relax Mesh” and then “Save Alignment” after the grid has been re-relaxed onto the mesh
184.
185.
187.
188.
189.
190.
191.
192.
193.
194.
195.
196.
203.
204.
Note: using a HD mesh with AlignMe will make the app run more slowly

## Slide 38
Get relaxed optode positions from AlignMe and visualize HD mesh + relaxed pad
% get relaxed optode positions from AlignMe
tposNew_HD = ds_HD.dO.tpos2_relaxed;
% visualize mesh with relaxed optodes
pM.reg=0;
PlotMeshSurface(visMeshHD,pM);Draw_Foci_191203(tposNew_HD,paramsFoci)
view(0,0) %posterior view
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
206.
207.
209.
210.
211.
212.

## Slide 39
Update pad info structure after running AlignMe
MRI Acquisition
Image Segmentation
Mesh Generation
Optode Placement
After the pad has been relaxed onto the HD mesh, we need to update the info structure with the new optode locations
info.optodes.spos3=tposNew_HD(1:Ns,:);
info.optodes.dpos3=tposNew_HD((Ns+1):end,:);
m=0;
for d=1:Nd
for s=1:Ns
m=m+1;
info.pairs.r3d(m)=norm(tposNew_HD(s,:)-tposNew_HD(d+Ns,:));
info.pairs.r3d(m+(Ns*Nd))=info.pairs.r3d(m);
end
end
% histogram of number of measurements per SD vs radius between SD pairs figure;histogram(info.pairs.r3d,1000);xlabel('R_S_D');ylabel('N_m_e_a_s');
% save pad file
gridname=padname;
save(['Pad',meshname,'_',gridname,'_', '.mat'],'info')
For both LD and HD meshes
216.
217.
218.
219.
220.
221.
222.
223.
224.
225.
227.
228.
230.
231.
232.
Nearest neighbors (left)
Lower radius between SD’s (right) correspond to lower nearest neighbors, i.e. NN1 and NN2

## Slide 40
Load in Additional Variables
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Please load: sfNIRS_NeuroDOT_workshop_additional_variables.mat
We will use these to visualize the HD mesh + relaxed optode positions and generate the sensitivity matrix
This file contains
  HD mesh
  Relaxed source/detector positions (standalone variable)
  Info file containing relaxed pad info
  “Gridname” variable
  Parameters

## Slide 41
Calculate the sensitivity (1 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Prepare mesh and optode information for NIRFAST, then visualize
gridname=padname;
mesh=PrepareMeshForNIRFAST(mesh,[meshname,'_',gridname],tposNew);
PlotMeshSurface(mesh,pM);Draw_Foci_191203(tposNew_HD, paramsFoci);
view([0,0]) %posterior view
m3=CutMesh(mesh,intersect(find(mesh.nodes(:,3)>0),find(mesh.nodes(:,1)>0)));
[Ia,Ib]=ismember(m3.nodes,mesh.nodes,'rows');Ib(Ib==0)=[];
m3.region=mesh.region(Ib);
PlotMeshSurface(m3,pM);Draw_Foci_191203(tposNew_HD, paramsFoci);
view([135,40]) %view mesh from behind with an angled top-down view
For both LD and HD meshes
1.
2.
236.
237.
238.
239.
245.
246.
247.
248.
249.
Visualize cap and mesh with portion cut out to make extra sure that optodes are not placed too far within the mesh

## Slide 42
Calculate the sensitivity (2 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
3. Set parameters for light modeling
flags.tag=[padname,'_on_',meshname,'_test'];
flags.gridname=gridname;
flags.meshname=meshname;
flags.head='info';
flags.info=infoT1;
flags.gthresh=1e-5;
flags.voxmm=2;
flags.labels.r1='csf';
flags.labels.r2='white';
flags.labels.r3='gray';
flags.labels.r4='bone';
flags.labels.r5='skin';
flags.op.lambda=[750,850];
flags.op.mua_skin=[0.0170,0.0190];
flags.op.mua_bone=[0.0116,0.0139];
flags.op.mua_csf=[0.0040,0.0040];
flags.op.mua_gray=[0.0180,0.0192];
flags.op.mua_white=[0.0167,0.0208];
flags.op.musp_skin=[0.74,0.64];
flags.op.musp_bone=[0.94,0.84];
flags.op.musp_csf=[0.3,0.3];
flags.op.musp_gray=[0.8359,0.6726];
flags.op.musp_white=[1.1908,1.0107];
flags.op.n_skin=[1.4,1.4];
flags.op.n_bone=[1.4,1.4];
flags.op.n_csf=[1.4,1.4];
flags.op.n_gray=[1.4,1.4];
flags.op.n_white=[1.4,1.4];
flags.srcnum=Ns;
flags.t4=affine_Subj2MNI;
flags.t4_target='MNI'; % string
flags.makeA=1; % don't make A, just make G
flags.Hz=0;
if flags.Hz, flags.tag = [flags.tag,'FD']; end
For both LD and HD meshes
255.
256.
257.
258.
259.
260.
261.
262.
263.
264.
265.
266.
267.
268.
269.
270.
271.
272.
273.
274.
275.
276.
277.
278.
279.
280.
281.
282.
283.
284.
285.
286.
287.
288.
Region labels
Absorption Coefficients for regions
Scattering Coefficients for regions
Index of refraction for regions
Voxelation resolution

## Slide 43
If makeAnirfaster is successful, output like the below screenshot should show up in your command window
Calculate the sensitivity (3 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
4. Calculate Green’s functions, voxelate, create sensitivity matrix, adjust bookkeeping
Ti=tic;[A,dim,Gsd]=makeAnirfaster(mesh,flags); % size(A)= [Nwl, Nmeas, Nvox]
disp(['<makeAnirfast took ',num2str(toc(Ti))])
%Gsd vs Rsd provides a simulated light fall-off figure;semilogy(info.pairs.r3d(1:(Ns*Nd)),Gsd,'*');
legend({num2str(flags.op.lambda')})
xlabel('R_S_D [mm]');ylabel('G_S_D');xlim([0,100])
For both LD and HD meshes
291.
292.
294.
295.
296.
297.

## Slide 44
Package Sensitivity Matrix with associated info
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
5. Arrange A to have the same measurement list structure as the data:
[Nwl,Nmeas,Nvox]=size(A);
A=reshape(permute(A,[2,1,3]),Nwl*Nmeas,Nvox);
6. Place spatial information about light model in info.tissue structure:
info.tissue.dim=dim;
info.tissue.affine=flags.t4;
info.tissue.infoT1=infoT1;
info.tissue.affine_target='MNI';
info.tissue.flags=flags;
save(['A_',flags.tag,'.mat'],'A','info','-v7.3’)
For both LD and HD meshes
301.
302.
305.
306.
307.
308.
309.
311.

## Slide 45
Validate the sensitivity (1 of 3)
Visualize an example measurement’s sensitivity using the A matrix containing measurements from SD’s with separations < 6cm. We’ll start with a single source detector pair and visualize the full sensitivity profile (FFR) on slide 48
%% Visualize aspects of sensitivity profile
t1=affine3d_img(mask,infoT1,dim,eye(4)); % put anatomical volume in dim space
keep=info.pairs.WL==2 & info.pairs.Src==1 & info.pairs.Det==6;
foo=squeeze(A(keep,:)); % Single meas pair
fooV=Good_Vox2vol(foo',dim);
fooV=fooV./max(fooV(:));
fooV=log10(1e2.*fooV); % top 2 o.o.m.
pA.PD=1;pA.Scale=2;pA.Th.P=0;pA.Th.N=-pA.Th.P;
PlotSlices(t1,dim,pA,fooV)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes
Single SD pair: source 1 detector 6
314.
315.
317.
318.
319.
320.
321.
322.
323.

## Slide 46
Validate the sensitivity (1 of 3)
Click on the activation sites seen on each slice to better see the sensitivity profile for measurements recorded from Source 2 and Detector 2
%% Visualize aspects of sensitivity profile
t1=affine3d_img(mask,infoT1,dim,eye(4)); % put anatomical volume in dim space
keep=info.pairs.WL==2 & info.pairs.Src==2 & info.pairs.Det==2;
foo=squeeze(A(keep,:)); % Single meas pair
fooV=Good_Vox2vol(foo',dim);
fooV=fooV./max(fooV(:));
fooV=log10(1e2.*fooV); % top 2 o.o.m.
pA.PD=1;pA.Scale=2;pA.Th.P=0;pA.Th.N=-pA.Th.P;
PlotSlices(t1,dim,pA,fooV)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes
Single SD pair: source 2 detector 2
314.
315.
317.
318.
319.
320.
321.
322.
323.

## Slide 47
Validate the sensitivity (3 of 3)
Visualize the flat field reconstruction
keep=(info.pairs.WL==2 & info.pairs.r2d<=40);
a=squeeze(A(keep,:));
iA=Tikhonov_invert_Amat(a,0.01,0.1);
iA=smooth_Amat(iA,dim,5); % use smoothing parameter of 5 here
ffr=makeFlatFieldRecon(a,iA);
fooV=Good_Vox2vol(ffr,dim);
fooV=fooV./max(fooV(:));
pA.PD=1;pA.Scale=1;pA.Th.P=1e-2;pA.Th.N=-pA.Th.P;
PlotSlices(t1,dim,pA,fooV)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes
337.
338.
339.
340.
341.
343.
344.
345.
346.

## Slide 48
Put it all together
As a final check, let’s visualize the alignment of the LD mesh, HD mesh, optode array, and cortical meshes all together
% View alignment
Cortical_mesh = load(['MNI164k_big.mat']);
Anat.CtxL = Cortical_mesh.MNIl;
Anat.CtxR = Cortical_mesh.MNIr;
pA0l=struct;
figure('Color','k','Position',[500,100,1050,1000])
pM1.fig_handle=gca;pM1.FaceColor=[1,1,1].*0.5;
pM1.SpecularStrength=0.25;pM1.DiffuseStrength=0.5;
PlotMeshSurface(meshHD,pM1) % HD mesh
pA0l.fig_handle=gca;
pA0l.FaceColor=[1,1,1].*0.5;pA0l.EdgeColor='none';
pA0l.AmbientStrength=0.25;pA0l.DiffuseStrength=0.25;
pA0l.SpecularStrength=0.025;
PlotMeshSurface(Anat.CtxL,pA0l) % Cortical Surfaces
PlotMeshSurface(Anat.CtxR,pA0l)
pA2.fig_handle=gca;
pA2.FaceAlpha=0;
pA2.EdgeAlpha=0.5;
pA2.EdgeColor=[1,1,1].*0.25;
PlotMeshSurface(meshLD,pA2) % LD mesh
set(gca,'Color','k')
view([90,0]) % display in lateral view
Draw_Foci_191203(cat(1,info.optodes.spos3,info.optodes.dpos3),paramsFoci) % Array
axis off
% For posterior view uncomment the following line
% view([0,0])
350.
351.
352.
353.
354.
355.
356.
357.
358.
359.
360.
361.
362.
363.
364.
365.
366.
367.
368.
369.
370.
371.
372.
373.
375.
376.

## Slide 49: Appendix: Generating info for Pad file (1 of 3)
The Pad file for a given grid design contains a structure variable info that contains the spatial and topological information for the array as follows:
  info.optodes
    Structure containing spatial coordinates for sources and detectors in array
    info.optodes.spos2 – 2D coordinates of each sources
    info.optodes.spos3 – 3D coordinates of each source
    info.optodes.dpos2 – 2D coordinates of each detector
    info.optodes.dpos3 – 3D coordinates of each detector
  info.pairs
    Table containing topological information for each source-detector measurement pair
    Each row of this table corresponds to a measurement and must be in the same order as the optical data for valid processing
    Generally, full list of sources and detectors is repeated for each wavelength/lambda.
    info.pairs.Src – source in measurement
    info.pairs.Det – detector in measurement
    info.pairs.NN – Nearest Neighbor class in array (e.g., info.pairs.NN=1 means this measurement is one of the closest in the grid design)
    info.pairs.WL – wavelength index of measurement
    info.pairs.lambda – actual wavelength of measurement (e.g., info.pairs.lambda-=750 is a 750 nm measurement)
    info.pairs.Mod – Modulation type of measurement (e.g., ‘CW’, or, ‘FD’)
    info.pairs.r2d – 2D source-detector distances
    info.pairs.r3d – 3D source-detector distances

## Slide 50: Appendix: Generating info for Pad file (2 of 3)
If you already have a grid design with 2D and 3D source and detector locations, saved as the following:
  grid.spos2
  grid.spos3
  grid.dpos2
  grid.dpos3
To generate the topology structure for your grid:
  dr=5; % Distance separating 'nearest neighbor' groups
  Rad=Grid2Radius_180824(grid,dr);
Once you have the grid and radius set up (as above), the info structure can be generated with the following:
  grid.name='Your_Grid_Name';
  params.lambda=[750,850]; % Wavelengths for your system
  params.Mod='CW'; % Modulation strategy
  [info]=Generate_Info_from_Grid_Rad(grid,Rad,params);
Then just save the Pad file for future use:
  save(['Pad_',info.optodes.CapName],'info')

## Slide 51: Appendix: Generating info for Pad file (3 of 3)
If instead you are designing a grid from scratch in style of our visual cap:
  Ndy=4; % Num detectors in y-direction
  Ndx=7; % Num detectors in x-direction
  NN1sep=13; % Closest source-detector Separation
  Srad=80; % Radius of spherical head model
  [grid,Rad]=Make_Sphere_Cap(Ndy,Ndx,NN1sep,Srad);
As described above,
  grid contains information on the positions of the sources and detectors.
  Rad contains information on the topology of the array.
Visualize the grid in 2D and 3D
  PlotSD(grid.spos2,grid.dpos2,'norm'); % Visualize 2D positions
  PlotSD(grid.spos3,grid.dpos3,'norm'); % Visualize 3D positions
Then set parameters, generate info file as on previous slide, and save Pad file for future use.

## Slide 52: That’s It (For Now)
Congratulations! You have finished the NeuroDOT Light Modeling Pipeline Tutorial for the basic 24x28 visual pad using AlignMe.
See other Sample Results Appendices for example visualizations of these same processing steps on the other Sample Data included in the toolbox.
Also, see the PreProcessing and Reconstruction tutorials to gain a deeper understanding of the effects of altering the processing parameters.
For further questions or more information, please consult the NeuroDOT User Manual and the various Appendices.
NeuroDOT Team:
  Adam Eggebrecht (aeggebre@wustl.edu)
  Emma Speh (espeh@wustl.edu)
  Ari Segel (ari@wustl.edu)
  Abby Magee (abigail.magee@wustl.edu)
  Weihao Fan (weihao.f@wustl.edu)