## Slide 1: NeuroDOT
Tutorial:
Generating a Light Model
32x32 Full Head Pad
with Hummingbird

## Slide 2
NeuroDOT Set-Up - Checklist
Before starting any tutorial, please make sure you’ve done the following
Download the appropriate versions of matlab, neuroDOT, and NIRFAST
  Matlab: 2020b
  NeuroDOT toolbox from NITRC
    https://www.nitrc.org/projects/neurodot/
  NIRFASTer-master
    https://github.com/nirfaster/NIRFASTer
Set your path so that all required folders and subfolders are included
  Folders w/subfolders to include
    NeuroDOT toolbox
    NIRFASTer-master
    Your personal directory
  Folders to include
    Output directory

## Slide 3
NeuroDOT Set-Up – Output Directory
Once you’ve set your path, the next thing you should do is set your output directory.
  The output directory must be located within your path
  If it isn’t already in your path, you must add it to your path
You can set your own path to an output directory in the line shown above
An output directory specifies the location where files will be saved
  MATLAB, by default, will save files to the current directory
The following line of code will change your current directory to the output directory that you just specified 
Now you can save outputs directly to your output directory
Example output directory

## Slide 4: Diffuse Optical Tomography
In diffuse optical tomography (DOT), arrays of optical sources and detectors are used to perform functional neuroimaging experiments on task-based or resting state modes of human brain function.
Analyses are broken into several pipelines:
  B: modeling of the tissue shape, optical property distribution, and source/detector locations
  C: modeling of the light emission, diffusion, and detection through the head
  D: preprocessing of the raw source-detector measurements
  E: reconstruction and spectroscopy of the preprocessed data and light model into a functional neuroimaging volume
  F: post-processing analysis of these results.

## Slide 5: Light Modeling Tutorials in NeuroDOT
NeuroDOT has several different tutorials that detail how to Generate a Light Model, each with its own purpose. They’re listed below to help you figure out which version is the best for your purposes.
Basic Light Modeling Tutorial with AlignMe
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_Pad_24x28_With_AlignMe
  Script: Basic_Light_Modeling_with_AlignMe_Tutorial
  Pad file: Pad_AdultV24x28.mat
Sparse Pad Tutorial with AlignMe
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_SparsePad_32x32_With_AlignMe
  Script: Basic_Light_Modeling_with_sparse_Pad_and_AlignMe_Tutorial
  Pad file: Pad_FullHead_32x32.mat
Light Modeling with a Subset of Measurements
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_With_Subset_of_Measurements_Pad_24x28_and_AlignMe
  Script: Light_Modeling_with_Subset_of_Measurements_and_AlignMe_Tutorial
  Pad file: Pad_AdultV24x28.mat
Split Pad Tutorial
  PPT: NeuroDOT_Tutorial_Split_Pad_Light_Modeling_With_LUMO_Adult_Temporal_Pad_and_AlignMe
  Script: Split_Pad_Light_Modeling_with_LUMO_Adult_Temporal_Pad_and_AlignMe_Tutorial
  Pad file: LUMO_adult_temporal_pad.mat
Basic Light Modeling with Hummingbird
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_Pad_24x28_With_Hummingbird
  Script: Basic_Light_Modeling_with_Hummingbird_Tutorial
  Pad file: Pad_AdultV24x28.mat
Sparse Pad Tutorial with Hummingbird
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_SparsePad_32x32_With_Hummingbird
  Script: Basic_Light_Modeling_with_sparse_Pad_and_Hummingbird_Tutorial
  Pad file: Pad_FullHead_32x32.mat
Light Modeling with Adult 96x92 Pad
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_Pad_Adult_96x92_Example
  Script: Basic_Light_Modeling_with_AlignMe_Tutorial_Pad_Adult_96x92
  Pad file: Pad_Adult_96x92.mat
  Note: this tutorial’s PowerPoint is in an older style

## Slide 6: Generating a Light Model
This tutorial will follow the Basic_Light_Modeling_with_sparse_Pad_and_Hummingbird_Tutorial.m Matlab file which can be found in the Documentation/Scripts folder.
The tutorial will:
  Generate a mesh from a segmented volume
  Place an array of sources and detectors on the mesh
  Set multiple parameters for the model
  Calculate the sensitivity A for this model
We will use a segmented volume, Segmented_MNI152nl_on_MNI111, found in Support_Files/Atlases, to generate the head model.
We will use the visual pad, Pad_FullHead_32x32.mat, found in Support_Files/Pad.

## Slide 7
NeuroDOT Flowchart for Light Model Generation
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
These steps are repeated for a low-density and high-density mesh

## Slide 8
Running the script – Author’s Note
How the author runs the script
  Highlight and execute each chunk of code
  Ex: highlight lines 64-69  right click  hit “Evaluate Selection”
Alternate way of running the script
  For each section (beige) hit ctrl+enter
  This isn’t how the author does it, but it will still work
Note: DO NOT HIT THE GREEN “RUN” BUTTON to run this script
MATLAB versions to use
  Optimal version: MATLAB 2020b
    The version that AlignMe was developed in
NIRFAST version to use
  NIRFASTer, the newest release of NIRFAST
  Found here: https://github.com/nirfaster/NIRFASTer
Example
section

## Slide 9
NeuroDOT Set-Up – In Script
The first section of code in this script will add the required toolboxes to your path and allow you to navigate to your desired output directory from within the script
Some parameters for visualizations made in the script are set up here as well
%Set Output Directory
%Change the path on the following line to whatever output directory you want to use
outputdir = ‘’; % put path to directory in between quotes
cd(outputdir)
% Does your pad have shorter 2D positions for visualizing the pad in 2D?
% i.e. does your pad have optode positions in an EEG style?
% if 2D positions are shorter, set style = 1. Otherwise, set style = 0
style = 1;
%Parameter Initialization
p.Cmap='jet'; p.Scale=5; p.Th.P=0; p.Th.N=-p.Th.P; p.PD=1; p.BG=[0,0,0];
pM.orientation='coord'; pM.Cmap.P='gray'; pM.EdgeColor='none';
pS = pM; pS = rmfield(pS, 'orientation'); % make copy of params, to be used with PlotSlices

## Slide 10
Running the script – Author’s Note
How the author runs the script
  Highlight and execute each chunk of code
  Ex: highlight lines 64-69  right click  hit “Evaluate Selection”
Alternate way of running the script
  For each section (beige) hit ctrl+enter
  This isn’t how the author does it, but it will still work
Note: DO NOT HIT THE GREEN “RUN” BUTTON to run this script
MATLAB versions to use
  Optimal version: MATLAB 2020b
    The version that AlignMe was developed in
NIRFAST version to use
  NIRFASTer, the newest release of NIRFAST
  Found here: https://github.com/nirfaster/NIRFASTer
Example
section

## Slide 11
NeuroDOT Set-Up – In Script
The first section of code in this script will add the required toolboxes to your path and allow you to navigate to your desired output directory from within the script
Some parameters for visualizations made in the script are set up here as well

## Slide 12
Load a Segmented volume and visualize
Note that PlotSlices.m generates an interactive visualization of the volume. Simply press ‘q’ or the middle mouse button to release the figure.
MRI Acquisition
Image Segmentation
5 – Scalp
4 – Skull
3 – Gray Matter
2 – White Matter
1 – CSF
0 - Background

## Slide 13
Generate a low-density mesh and visualize
The low-density mesh will be used to efficiently align the optical array to its position on the head.
MRI Acquisition
Image Segmentation
LD Mesh Generation
“coordinate space”
The origin is in the center of volume
A view inside

## Slide 14
Place the optodes on the LD mesh
Load in a Pad file for a grid design of interest, here we use the AdultV24x28 array.
Note: This Pad file is a starting point and may not align with the mesh right away.
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
2D positions
3D positions

## Slide 15
Check Optode Separation against Nearest Neighbors
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
See, there are measurements for NN3, 4, and 5. This is what we want to correct
Before correction
After Correction
Now, only NN1 and NN2 remain
Note: The “before correction” plot is not generated in the MATLAB tutorial script
Nearest Neighbor vs Optode Separation scatterplot is a useful tool for catching potential issues for spring relaxation before they occur.
  Visualize info.pairs.NN agains info.pairs.r3d. If the plot contains NN3, NN4, etc. Correct the info structure so that measurements < 6cm have only NN1 and NN2

## Slide 16
Place the optodes on the LD mesh
Adjust array position via the Hummingbird App by altering parameters and re-visualizing mesh and array:
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
First, reset the grid position
  Select: “to Behind the Head” from the drop down
  The view will then update
Then, choose a view that will best show the location of the array relative to the mesh
  XY view is great for this
Input new values for parameters and Update Figure
  dx, dy, dz: translate the array
  dxTh, dyTh, dzTh: rotate the array
  Scale: size the array
  flipX, flipY, flipZ: flip the array
Repeat until array is ~5mm off the head for the whole pad but is lined up with the mesh. After following the upcoming tutorial, the mesh and grid should look like the image on the right.
Note: The following 15 slides are an in-depth tutorial on how to use the Hummingbird app

## Slide 17: Hummingbird Tutorial 1
Run the following code to create the DataStorage class and open the Hummingbird app
Default view on Hummingbird launch

## Slide 18: Hummingbird Tutorial 2
Change the view to fully view the pad
  The pad is the grid comprised of red and blue spheres; these denote optodes
    Red = Source
    Blue = Detector
  Some optodes might start within the mesh
Switch to XY view to start

## Slide 19: Hummingbird Tutorial 3
Change the view to fully view the pad
  The pad is the grid comprised of red and blue spheres; these denote optodes
    Red = Source
    Blue = Detector
  Some optodes might start within the mesh
Now try YZ view
  Note that some optodes aren’t visible right now because they’re inside the mesh
You can also manually set the view by editing the numbers
  Set the first value to -90
  Then hit “Update view”

## Slide 20: Hummingbird Tutorial 4
You can also view source and detector (SD) numbers by hitting the “View SD Spheres/Labels” button
Button changes after clicked
  View SD Labels
  View SD Spheres

## Slide 21: Hummingbird Tutorial 5
Back to XZ view
Set scale to 1.1
Hit “Update Figure” whenever you update a field in the Modify Parameters box
We want the optodes to be around 5-10mm AWAY from the head so that they can relax back on to it
Double check that all the optodes are away from the head by changing views

## Slide 22: Hummingbird Tutorial 6
Once you’ve verified that the pad position is correct and that optodes are away from the head, hit the “Relax Mesh” button
A pop-up box will then open, hit “Done”
The below two figures will open and start to update

## Slide 23: Hummingbird Tutorial 8
Once optodes are relaxed onto the mesh, a dialogue box will show up on the Hummingbird app
Hit “OK” to close it

## Slide 24: Hummingbird Tutorial 9
To save the new optode locations to your MATLAB workspace, hit the “Finalize” button
This will also close Hummingbird
Note: on some versions of MATLAB, the finalize button will not be visible
  Simply close Hummingbird and resume the script at the next line, the relaxed optodes will be saved to the DataStorage object and you will be able to access them
Congrats!! You’ve finished the Hummingbird tutorial!

## Slide 25
Generate smaller mask based on optode locations
Here, we get our relaxed grid position and visualize it on the LD mesh. Then, we scoop out the part of the mesh that the grid is sitting on. This scooped section will is the basis for the high-density mesh that we are about to create
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation

## Slide 26
Generate a high-density mesh and visualize
Take the scooped section of the LD mesh and turn it into a HD mesh
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Top View
Bottom View
Bottom view helps to make sure the mesh is totally hollow, with no connected strands on the inside

## Slide 27
Generate a high-density mesh and visualize
Notes and possible errors
Files to delete if File I/O Error occurs
Pop-up box for File I/O Error
Possible Errors in HD mesh generation
File I/O Error:
  Solution: delete all mesh files in your output directory, change the meshname variable to a new name and re-run
  Ex: ‘_HD_Mesh1’  ‘_HD_Mesh0’
  Make sure you delete these files from the file manager, and NOT within matlab
  If the above doesn’t work, create an entirely new output directory, make that directory your current folder, then re-run lines 149-168
Poor mesh quality
  Occurs if number of nodes in mesh is greater than 1,000,000
  This is why we make smaller HD meshes and tune params.cell size
Parameters
Cell_size = size of internal tetrahedra
  This can be changed to ensure that the number of nodes in the mesh does not exceed 999,999
  For HD meshes, this is typically set to 1.0 for visual pad, and 2.0 for larger volume meshes i.e. full head
Facet_distance = error tolerance at boundary for node positions
  Typically set to 2.0 for HD meshes
Facet_size = size of surface triangles of mesh
  Want to keep this as small as possible, as we want the surface to be highly faceted
  Typically set to 0.8 for HD meshes

## Slide 28
Place the optodes on the HD mesh
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Optode Placement
This time, since the HD mesh is in the same space that the LD mesh was, the grid will already be lined up and sitting on the mesh when hummingbird starts. You’ll need to move it away from the mesh and re-relax the grid back onto the mesh
Change the view to “XY View”
Move the grid off the mesh by setting dy = -10
Adjust the scale to 1.1 (this is the scale we used for the LD mesh too)
Hit “Update Figure”
  Hummingbird’s GUI should now look like the image below
Hit “Relax Mesh” and “Finalize” after the grid has been re-relaxed onto the mesh

## Slide 29
Update pad info structure after running Hummingbird
MRI Acquisition
Image Segmentation
Mesh Generation
Optode Placement
After the pad has been relaxed onto the HD mesh, we need to update the info structure with the new optode locations
For both LD and HD meshes

## Slide 30
Calculate the sensitivity (1 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Prepare mesh and optode information for NIRFAST, then visualize
Visualize cap and mesh with portion cut out to make extra sure that optodes are not placed too far within the mesh
For both LD and HD meshes
1.
2.

## Slide 31
Calculate the sensitivity (2 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
3. Set parameters for light modeling
For both LD and HD meshes
Region labels
Absorption Coefficients for regions
Scattering Coefficients for regions
Index of refraction for regions
Voxelation resolution

## Slide 32
Calculate the sensitivity (3 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
4. Calculate Green’s functions, voxelate, create sensitivity matrix, adjust bookkeeping
If makeAnirfast is successful, info like the below screenshot should show up in your command window
For both LD and HD meshes

## Slide 33
Package Sensitivity Matrix with associated info
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
5. Arrange A to have the same measurement list structure as the data:
6. Place spatial information about light model in infoHD.tissue structure:
For both LD and HD meshes

## Slide 34
Package Sensitivity Matrix with associated info
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
6. Package and save A where all source-detector separations are less than 5cm
For sparse pads, only measurements from source-detector (SD) pairs that have a separation distance within 5cm are generally used. The Gsd vs Rsd scatterplot is a helpful tool for visualizing what measurements meet the 5cm cutoff. A red box has been placed around the measurements within the 5cm cutoff.
For both LD and HD meshes
Gsd vs Rsd

## Slide 35
Validate the sensitivity (1 of 2)
Visualize an example measurement’s sensitivity
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes
NN1: source 1 detector 1

## Slide 36
Validate the sensitivity (2 of 2)
Visualize the flat field reconstruction
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes

## Slide 37
Put it all together
As a final check, let’s visualize the alignment of the LD mesh, HD mesh, optode array, and cortical meshes all together
Lateral View
Dorsal View
Posterior View

## Slide 38: That’s It (For Now)
Congratulations! You have finished the NeuroDOT Light Modeling Pipeline Tutorial for the 32x32 Full Head pad using Hummingbird.
See other Sample Results Appendices for example visualizations of these same processing steps on the other Sample Data included in the toolbox.
Also, see the PreProcessing and Reconstruction tutorials to gain a deeper understanding of the effects of altering the processing parameters.
NeuroDOT Team:
  Adam Eggebrecht (aeggebre@wustl.edu)
  Emma Speh (espeh@wustl.edu)
  Ari Segel (ari@wustl.edu)
  Abby Magee (abigail.magee@wustl.edu)
  Weihao Fan (weihao.f@wustl.edu)