## Slide 1: NeuroDOT
Tutorial:
Generating a Light Model
With a Split Pad

## Slide 2
NeuroDOT Set-Up - Checklist
Before starting any tutorial, please make sure you’ve done the following
Download the appropriate versions of matlab, neuroDOT, and NIRFAST
  Matlab: 2020b
  NeuroDOT toolbox from NITRC
    https://www.nitrc.org/projects/neurodot/
  NIRFASTer-master
    https://github.com/nirfaster/NIRFASTer
Set your path so that all required folders and subfolders are included
  Folders w/subfolders to include
    NeuroDOT toolbox
    NIRFASTer-master
    Your personal directory
  Folders to include
    Output directory

## Slide 3
NeuroDOT Set-Up – Output Directory
Once you’ve set your path, the next thing you should do is set your output directory.
  The output directory must be located within your path
  If it isn’t already in your path, you must add it to your path
You can set your own path to an output directory in the line shown above
An output directory specifies the location where files will be saved
  MATLAB, by default, will save files to the current directory
The following line of code will change your current directory to the output directory that you just specified 
Now you can save outputs directly to your output directory
Example output directory

## Slide 4: Diffuse Optical Tomography
In diffuse optical tomography (DOT), arrays of optical sources and detectors are used to perform functional neuroimaging experiments on task-based or resting state modes of human brain function.
Analyses are broken into several pipelines:
  B: modeling of the tissue shape, optical property distribution, and source/detector locations
  C: modeling of the light emission, diffusion, and detection through the head
  D: preprocessing of the raw source-detector measurements
  E: reconstruction and spectroscopy of the preprocessed data and light model into a functional neuroimaging volume
  F: post-processing analysis of these results.

## Slide 5: Light Modeling Tutorials in NeuroDOT
NeuroDOT has several different tutorials that detail how to Generate a Light Model, each with its own purpose. They’re listed below to help you figure out which version is the best for your purposes.
Basic Light Modeling Tutorial with AlignMe
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_Pad_24x28_With_AlignMe
  Script: Basic_Light_Modeling_with_AlignMe_Tutorial
  Pad file: Pad_AdultV24x28.mat
Sparse Pad Tutorial with AlignMe
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_SparsePad_32x32_With_AlignMe
  Script: Basic_Light_Modeling_with_sparse_Pad_and_AlignMe_Tutorial
  Pad file: Pad_FullHead_32x32.mat
Light Modeling with a Subset of Measurements
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_With_Subset_of_Measurements_Pad_24x28_and_AlignMe
  Script: Light_Modeling_with_Subset_of_Measurements_and_AlignMe_Tutorial
  Pad file: Pad_AdultV24x28.mat
Split Pad Tutorial
  PPT: NeuroDOT_Tutorial_Split_Pad_Light_Modeling_With_LUMO_Adult_Temporal_Pad_and_AlignMe
  Script: Split_Pad_Light_Modeling_with_LUMO_Adult_Temporal_Pad_and_AlignMe_Tutorial
  Pad file: LUMO_adult_temporal_pad.mat
Basic Light Modeling with Hummingbird
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_Pad_24x28_With_Hummingbird
  Script: Basic_Light_Modeling_with_Hummingbird_Tutorial
  Pad file: Pad_AdultV24x28.mat
Sparse Pad Tutorial with Hummingbird
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_SparsePad_32x32_With_Hummingbird
  Script: Basic_Light_Modeling_with_sparse_Pad_and_Hummingbird_Tutorial
  Pad file: Pad_FullHead_32x32.mat
Light Modeling with Adult 96x92 Pad
  PPT: NeuroDOT_Tutorial_Generating_a_Light_Model_Pad_Adult_96x92_Example
  Script: Basic_Light_Modeling_with_AlignMe_Tutorial_Pad_Adult_96x92
  Pad file: Pad_Adult_96x92.mat
  Note: this tutorial’s PowerPoint is in an older style

## Slide 6: Generating a Light Model – Split Pad Tutorial
This tutorial will follow the Split_Pad_Light_Modeling_with_LUMO_Adult_Temporal_Pad_and_AlignMe_Tutorial.m Matlab file which can be found in the Documentation/Scripts folder.
The tutorial will:
  Generate a mesh from a segmented volume
  Place an array of sources and detectors on the mesh using the AlignMe App
  Set multiple parameters for the model
  Calculate the sensitivity A for this model for only measurements with a source detector separation of 4cm or less.
We will use a segmented volume, Segmented_MNI152nl_on_MNI111_nifti, found in Support_Files/Atlases, to generate the head model.
We will use the visual pad, LUMO_adult_temporal_pad.mat, found in Support_Files/Pad.

## Slide 7
NeuroDOT Flowchart for Light Model Generation
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
These steps are repeated for a low-density and high-density mesh

## Slide 8
Running the script – Author’s Note
How the author runs the script
  Highlight and execute each chunk of code
  Ex: highlight lines 64-69  right click  hit “Evaluate Selection”
Alternate way of running the script
  For each section (beige) hit ctrl+enter
  This isn’t how the author does it, but it will still work
Note: DO NOT HIT THE GREEN “RUN” BUTTON to run this script
MATLAB versions to use
  Optimal version: MATLAB 2020b
    The version that AlignMe was developed in
NIRFAST version to use
  NIRFASTer, the newest release of NIRFAST
  Found here: https://github.com/nirfaster/NIRFASTer
Example
section

## Slide 9
NeuroDOT Set-Up – In Script
The first section of code in this script will have you set your output directory and allow you to navigate to your desired output directory from within the script
Some parameters for visualizations made in the script are set up here as well

## Slide 10
Load a Segmented volume and visualize
MRI Acquisition
Image Segmentation
5 – Scalp
4 – Skull
3 – Gray Matter
2 – White Matter
1 – CSF
0 - Background
Note that PlotSlices.m generates an interactive visualization of the volume. Simply press ‘q’ or the middle mouse button to release the figure.

## Slide 11
Generate a low-density mesh and visualize
MRI Acquisition
Image Segmentation
LD Mesh Generation
“coordinate space”
The origin is in the center of volume
A view inside
The low-density mesh will be used to efficiently align the optical array to its position on the head
Mesh generation parameters of note
  facet_distance: error tolerance at boundary for node positions
  facet_size: size of surface triangles
  cell_size: size of internal tetrahedra
  Mode: controls density of mesh, 1 for low density, 0 for high density

## Slide 12
Load the Pad File
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
Load in a Pad file for a grid design of interest, here we use the LUMO Adult Temporal Pad
Since this array has sections that cover the left and right sides of the head, we’ll need to split the pad up into those sections – the 2D layout is very helpful for this
  Find the greatest source and detector numbers on the side of the pad that has S1 and D1. We’ll use these as indices to split the pad

## Slide 13
Make Right Side Pad
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
We’ll now create a pad file for the right hand side optodes
  Place the RHS optode positions in a structure
  Run Generate_pad_from_grid
  Visualize in 2D and 3D to make sure everything looks correct

## Slide 14
Make Left Side Pad
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
Then we’ll do the same thing for the left hand side optodes

## Slide 15
AlignMe Functionality
AlignMe allows for the user to scale the mesh to participant specific space by using
  Participant Head Circumference
  Participant Z line distance (nasion to inion)
These are contained within the Atlas Scaling box
Additionally, to facilitate easier alignment for participant specific situations, photos of the capfit can be utilized, and will show up where the Brain Light Lab logo is located on the GUI
  Rotate CW, Rotate CCW: rotates photo
  Previous, Next: scroll through available photos
Participant photos can be loaded in by setting the path to the directory containing photos. AlignMe will pull all files of the designated image type
  Path to photos: ds.dI.photopath
  Type of image: ds.dI.imageType
Change the fields in the Array Alignment box to position the array close to the head
  Then press the “relax array” button to allow the array to relax onto the surface of the head. The GUI should then look like the figure above
  Finally, use the “Save Alignment” button to save the relaxed optode positions to the workspace
NeuroDOT has an in depth walk through of how to use AlignMe in a standalone tutorial titled: NeuroDOT_Tutorial_AlignMe
  The tutorial can be found in the Documentation/Scripts folder of the NeuroDOT toolbox

## Slide 16
Relax Right Side Optodes
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
Now we’ll use AlignMe to position the right hand side pad
Use the fields in the Array Alignment box outlined in red to position the array
Use the view buttons outlined in green to change the view of the brain

## Slide 17
Relax Left Side Optodes
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
Now we’ll use AlignMe to position the left hand side pad
Use dx to move the array to the left side of the mesh
You should keep the same values for dy and dz as the right hand side pad
  Remember, these are two halves of the same pad

## Slide 18
Generate a high-density mesh and visualize
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Now, we’ll create a high-density mesh. HD meshes contain a large number of nodes – close to 1,000,000 nodes but never more – in addition to some anatomical information
Notice the difference between the parameters used in LD mesh generation on slide 11 and those used here
  Facet_distance is decreased
  Facet_size is decreased, leading to more triangles that make up the surface of the mesh
  Cell_size is decreased, there are more internal tetrahedra in HD meshes than LD meshes
Notes: NOTE: WE ARE NOT HAVING THEM RUN THESE SLIDES

## Slide 19
Possible errors during HD mesh generation
Files to delete if File I/O Error occurs
Pop-up box for File I/O Error
File I/O Error:
  Solution: delete all mesh files in your output directory, change the meshname variable to a new name and re-run
  Ex: ‘_HD_Mesh1’  ‘_HD_Mesh0’
  Make sure you delete these files from the file manager, and NOT within matlab
  If the above doesn’t work, create an entirely new output directory, make that directory your current folder, then re-run lines 149-168
Poor mesh quality
  Occurs if number of nodes in mesh is greater than 1,000,000
  This is why we make smaller HD meshes and tune params.cell size

## Slide 20
Place the optodes on the HD mesh
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
Optode Placement
This time, since the HD mesh is in the same space that the LD mesh was, the grid will already be lined up and sitting on the mesh when AlignMe starts. You’ll need to move it away from the mesh and re-relax the grid back onto the mesh
Change the view to “XY View”
Adjust the scale to 1.1 (this is the scale we used for the LD mesh too)
Move the grid off the mesh by setting dx = 10
  The AlignMe GUI should now look like the image below
Hit “Relax Mesh” and then “Save Alignment” after the grid has been re-relaxed onto the mesh
Repeat steps 1-5 for the Left hand side pad, but set dx = -10
Note: using a HD mesh with AlignMe will make the app run more slowly

## Slide 21
Get relaxed optode positions from AlignMe and visualize HD mesh + relaxed pad
MRI Acquisition
Image Segmentation
LD Mesh Generation
Optode Placement
HD Mesh Generation
At the end of each AlignMe section, the relaxed optodes will be visualized on the HD mesh
Left
Right

## Slide 22
Put left and right pads back together
MRI Acquisition
Image Segmentation
Mesh Generation
Optode Placement
For both LD and HD meshes
Now we’ll take the relaxed optode positions from both the left and right pads and put them together
After that, we’ll run two checks to check our work
  Make sure the total number of sources and detectors match the numbers specified in the original pad file
  Visualize all relaxed optodes on the mesh

## Slide 23
Update pad info structure after running AlignMe
MRI Acquisition
Image Segmentation
Mesh Generation
Optode Placement
For both LD and HD meshes
After the pad has been relaxed onto the HD mesh, we need to update the info structure with the new optode locations
We’ll use those optode locations to re-calculate the 3D source-detector separations for every source-detector pair
After that we’ll make a histogram showing the relationship of measurements per separation distance

## Slide 24
Calculate the sensitivity (1 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Prepare mesh and optode information for NIRFAST, then visualize
Visualize cap and mesh with portion cut out to make extra sure that optodes are not placed too far within the mesh
For both LD and HD meshes
1.
2.

## Slide 25
Calculate the sensitivity (2 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
3. Set parameters for light modeling
For both LD and HD meshes
Region labels
Absorption Coefficients for regions
Scattering Coefficients for regions
Index of refraction for regions
Voxelation resolution

## Slide 26
If makeAnirfaster is successful, output like the below screenshot should show up in your command window
Calculate the sensitivity (3 of 3)
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
4. Calculate Green’s functions, voxelate, create sensitivity matrix, adjust bookkeeping
For both LD and HD meshes

## Slide 27
Package Sensitivity Matrix with associated info
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
5. Reshape A so it’s the following size: Nwl*Nmeas X Nvox
6. Place spatial information about light model in info.tissue structure
For both LD and HD meshes

## Slide 28
Validate the sensitivity (1 of 3)
Let’s visualize the sensitivity of the A matrix. We’ll start with a single source detector pair (Source 1, Detector 1) and then visualize the full sensitivity profile (FFR) on slide 28
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes
Single SD pair: source 1 detector 1

## Slide 29
Validate the sensitivity (2 of 3)
Click on the areas that light up on each slice to better see the sensitivity profile for measurements recorded from Source 20 and Detector 25
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes
Single SD pair: source 20 detector 25

## Slide 30
Validate the sensitivity (3 of 3)
Visualize the flat field reconstruction
MRI Acquisition
Image Segmentation
Calculate Sensitivity
Mesh Generation
Optode Placement
Validate Sensitivity
For both LD and HD meshes

## Slide 31
Put it all together
As a final check, let’s visualize the alignment of the LD mesh, optode array, and cortical meshes all together
We are not visualizing the HD mesh because for this array, the HD mesh is the same size as the LD mesh

## Slide 32: That’s It (For Now)
Congratulations! You have finished the NeuroDOT Split Pad Light Modeling Pipeline Tutorial for the LUMO Adult Temporal pad using AlignMe.
See other Sample Results Appendices for example visualizations of these same processing steps on the other Sample Data included in the toolbox.
Also, see the PreProcessing and Reconstruction tutorials to gain a deeper understanding of the effects of altering the processing parameters.
NeuroDOT Team:
  Adam Eggebrecht (aeggebre@wustl.edu)
  Emma Speh (espeh@wustl.edu)
  Ari Segel (ari@wustl.edu)
  Abby Magee (abigail.magee@wustl.edu)
  Weihao Fan (weihao.f@wustl.edu)